## Slide 1
Face Landmark Tracking For Better ROI
Peak Finding Algorithm For Various Type of Signals
Rafsan Jany
Korean Convergence Medical Science

## Slide 2
Face ROI For HR Signals
ROI | HR Signal (RGB) | RR Signal (RGB) | Remarks | Key References
Forehead | Excellent | Poor | High perfusion, low motion, stable surface. Best for HR. | De Haan & Jeanne (2013), Li et al. (2014), Poh et al. (2010)
Left Cheek | Excellent | Moderate | Good perfusion, some respiratory modulation possible. | Verkruysse et al. (2008), Kumar et al. (2015), Yu et al. (2019)
Right Cheek | Excellent | Moderate | Same as left cheek — combined cheeks often used for better signal. | Kumar et al. (2015), McDuff et al. (2014), Yu et al. (2019)

## Slide 3
Face ROI For RR Signals
ROI | HR Signal (RGB) | RR Signal (RGB) | Remarks | Key References
Nostrils / Tip | Weak | Best | Visible subtle motion due to breathing. High RR signal quality. | Tarassenko et al. (2014), Harford et al. (2019), Nitzan et al. (2008)
Upper Lip | Weak | Excellent | Detects thermal and reflectance change with exhalation. Often used in RR estimation. | Harford et al. (2019), Wu et al. (2012)
Mouth Region | Noisy | Excellent | Airflow and soft tissue motion influence reflectance; helpful in RR extraction. | Yu et al. (2019), Tulyakov et al. (2016)
Chin | Poor | Poor | High motion, low signal strength. Typically avoided. | Chen & McDuff (2018), Kumar et al. (2015)

## Slide 4
ROI For HR And RR
Forehead
Right Cheek
Left Cheek
Nose Bridge
Nose Tip
Upper lip
Mouth Region

## Slide 5
Getting ROI For HR Using Mediapipe

## Slide 6
Getting ROI For HR Using Mediapipe

## Slide 7
ROI Evaluation
Gen | ICA MAE | FFT MAE
Full Face | 21.44 | 31.52
Static Bounding Box | 16.73 | 25.11
Dynamic Landmark based ROI | 12.34 | 18.61

## Slide 8
Peak Finding Algorithms
Category | Algorithm Name | Open-Source Availability | Reference
Derivative-Based | Zero-Crossing | SciPy | Rabiner, 1975 [Proc. IEEE]
 | Second Derivative Zero-Crossing | ❌ | Rabiner, 1975 [Proc. IEEE]
 | Pan-Tompkins Algorithm | biosppy, NeuroKit2 | Pan & Tompkins, 1985 [IEEE T-BME]
 | Squared Derivative Method | ❌ | Akin, EMBC, 2014
 | Differential Zero-Crossing | ❌ | Cardoso & Araujo, 2017 [IEEE Sensors]
Local Maxima | Standard Local Maxima (find_peaks) | SciPy | Virtanen et al., 2020 [SciPy]
 | Sliding Window Maxima | ❌ | Custom window-based method
 | Peak-Valley Detector | ❌ | Custom dual-polarity method
Envelope-Based | Hilbert Envelope Peaks | biosppy | Boashash, 1992 [Proc. IEEE]
 | Moving Average Envelope | ❌ | Sun et al., 2023 [Electronics]
Transform-Based | Autocorrelation-Based Peaks | NeuroKit2 | Oppenheim & Schafer [DSP Textbook]
 | Continuous Wavelet Transform (CWT) Peaks | PyWavelets, biosppy | Torrence & Compo, 1998 [BAMS]
 | FFT Envelope Peaks | ❌ | Oppenheim & Schafer [DSP Textbook]
Statistical/Energy | Short-Term Energy Peaks | ❌ | Oppenheim & Schafer [DSP Textbook]
 | RMS Energy Peaks | ❌ | Oppenheim & Schafer [DSP Textbook]
 | Spectral Entropy Gradient | ❌ | Novel method

## Slide 9
Peak Finding Algorithms
Category | Algorithm Name | Open-Source Availability | Reference
Gradient/SSF-Based | Slope Sum Function (SSF) | biosppy | Li et al., 2014 [IEEE TIP]
 | First-Order Difference Histogram | ❌ | Novel histogram method
 | Integrated Slope Method | ❌ | Li et al., 2014 [IEEE TIP]
AMPD & Variants | AMPD (Automatic Multiscale Peak Detection) | peakutils, NeuroKit2 | Scholkmann et al., 2012 [Algorithms]
 | Modified AMPD | ❌ | Scholkmann et al., 2012 [Algorithms]
 | Fast AMPD | ❌ | Scholkmann et al., 2012 [Algorithms]
 | FAMD (Frequency-Adaptive Multiscale) | ❌ | Scholkmann et al., 2012 [Algorithms]
Threshold-Based | Static Thresholding | basic Python | Classic threshold method
 | Adaptive Thresholding | NeuroKit2 | Zigel et al., 1998 [IEEE T-BME]
 | Max-in-Window Thresholding | ❌ | Custom implementation
Filtering-Based | Bandpass + Local Maxima | biosppy | Oppenheim & Schafer [DSP Textbook]
 | Savitzky–Golay Peak Detection | SciPy | Savitzky & Golay, 1964 [Anal. Chem.]
 | Morphological Gradient | skimage, SciPy | Gonzalez & Woods [Digital Image Processing]
 | Gaussian Smoothing Peaks | SciPy | Standard Gaussian filtering
 | Median Filtered Peaks | SciPy | Median filter-based smoothing

## Slide 10
Peak Finding Algorithms
Peak Finding Agorithm
Local Maxima
Envelope-Based
Transform-Based
Statistical/Energy
Gradient/SSF-Based
AMPD
Threshold-Based
Filtering-Based
Pan-Tompkins
Algorithm
Squared Derivative
Method
Differential
Zero-Crossing
Standard Local
Maxima
Sliding Window
Maxima
Peak-Valley Detector
Hilbert Envelope Peaks
Moving Average
Envelope
Autocorrelation-Based Peaks
Continuous Wavelet Transform (CWT) Peaks
FFT Envelope Peaks
Short-Term Energy
Peaks
RMS Energy Peaks
Spectral Entropy
Gradient
Slope Sum Function
(SSF)
First-Order Difference Histogram
Integrated Slope Method
Derivative-Based
Zero-Crossing
Second Derivative
Zero-Crossing
Automatic Multiscale Peak Detection
Modified AMPD
Fast AMPD
FAMD (Frequency-Adaptive Multiscale)
Statistical Threshold
Adaptive Threshold
Max Threshold
Morphological
Gradient
Gaussian Smoothing
Peaks
Median Filtered Peaks
Bandpass +
Local Maxima
Savitzky–Golay Peak
Detection

## Slide 11
Peak Finding Algorithms

## Slide 12
Peak Finding Algorithms

## Slide 13
Zero-Crossing Algorithms
Mechanism
Detects where the signal (or its derivative)
crosses zero, indicating slope changes
likely to correspond to peaks or troughs.
Advantages:
Simple and fast
No parameters needed
Suitable for regular oscillatory signals
Graph Explanation:
Captures rhythmic structure well but may
miss sharper peaks or falsely detect in flat
regions (near frames 90, 180).
Derivative-Based

## Slide 14
Second Derivative Zero-Crossing
Mechanism
Detects points where the second derivative crosses zero, indicating curvature changes (inflection points) often associated with peaks or valleys.
Advantages:
Highlights turning points with strong curvature
More sensitive than first derivative methods
Graph Explanation:
Numerous red dots appear near bends and curve transitions, including shoulders and minor fluctuations.
Tends to over-detect, especially in noisy or flat segments (frame 80–120), identifying many minor inflections.
Derivative-Based

## Slide 15
Pan-Tompkins
Mechanism
Applies bandpass filtering, derivative, squaring, and moving window integration to detect sharp changes, originally designed for
QRS complex in ECG
Advantages:
Highly selective for sharp, prominent peaks
Robust against noise and baseline drift
Graph Explanation:
Only one peak detected at the sharp negative spike near frame 5.
Method misses the regular rhythmic peaks because it targets sharp transients, not sinusoidal patterns typical in green channel rPPG signals.
Derivative-Based

## Slide 16
Squared Derivative
Mechanism
Computes the first derivative and squares it to emphasize sharp transitions. Peaks are identified where this energy exceeds a local threshold.
Advantages:
Highlights strong signal changes
Filters out low-amplitude noise
Enhances prominent rising/falling edges
Graph Explanation:
Peaks (red dots) align with steep transitions, especially where amplitude changes rapidly (frames ~5, 140, 240).
Misses smoother oscillations and favors sharp edges, leading to fewer detected peaks in the middle rhythmic section.
Derivative-Based

## Slide 17
Standard Local Maxima
Mechanism
Identifies local maxima by comparing each point to its neighbors, with options to set minimum distance, height, or prominence thresholds.
Advantages:
Widely used and customizable
Robust against noise via threshold tuning
Available in standard libraries (SciPy)
Graph Explanation:
Detected peaks (red dots) align with the most prominent maxima (frames ~40, 140, 170).
Conservative detection, only strong peaks are selected, ignoring smaller oscillations in flatter regions.
Local Maxima

## Slide 18
Sliding Window Maxima
Mechanism
Uses a moving window to identify local maxima within each segment. The center point is marked as a peak if it’s the highest in the window.
Advantages:
Simple and interpretable
Effective for signals with regular spacing
Window size controls sensitivity
Graph Explanation:
Red dots mark local peaks captured within each sliding window (frames ~35, 145, 180).
Balances between sharp and smooth peaks, giving good coverage across the signal.
Some closely spaced peaks are missed if window is too wide.
Local Maxima

## Slide 19
Peak-Valley Detector
Mechanism
Identifies alternating local maxima and minima by comparing each point to its neighbors. A peak must be followed by a valley and vice versa.
Advantages:
Detects complete oscillation cycles
Works well on periodic signals
Simple and interpretable
Graph Explanation:
Red dots align with major peaks in the rhythmic pattern (e.g., frames ~40, 140, 180).
Captures peak-to-valley alternations, providing consistent detection in regular waveforms.
Local Maxima

## Slide 20
Hilbert Envelope Peaks
Mechanism
Applies the Hilbert Transform to generate an analytic signal and compute its amplitude envelope. Peaks are then detected on this smoothed envelope.
Advantages:
Good at handling noisy or irregular signals
Highlights overall signal energy trend
Smoothens rapid fluctuations
Graph Explanation:
Peaks (red dots) are selected from the smoothed envelope rather than the raw waveform. Many small oscillations are ignored, favoring broader peaks (frame ~40, 130, 170).Less responsive to fine, high-frequency changes.
Envelope-Based

## Slide 21
Moving Average Envelope
Mechanism
Applies a moving average filter to smooth the signal, forming an envelope. Peaks are identified where the smoothed envelope rises and falls.
Advantages:
Reduces high-frequency noise
Emphasizes broader oscillatory trends
Customizable smoothness via window size
Graph Explanation:
Red points show detected peaks from the smoothed signal envelope.
Focuses on large-scale oscillations, missing finer peaks (frames ~130, 240).
Envelope-Based

## Slide 22
Autocorrelation-Based Peaks
Mechanism
Computes the autocorrelation of the signal to highlight periodicity.
Peaks are identified at regular time lags indicating signal repetition.
Advantages:
Excellent for periodic signal detection (heart rate).
Robust to noise and phase shifts.
Captures global rhythm, not local spikes.
Graph Explanation:
Red points correspond to dominant periodic positions from autocorrelation analysis.
Some peaks are missed due to focus on repeating intervals, not local maxima.
Useful in estimating average heart rate, but not ideal for all beat detection.
Transform-Based

## Slide 23
FFT Envelope Peaks
Mechanism
Applies the Fast Fourier Transform (FFT) to the signal to identify the dominant frequency components.
Peaks are detected in the envelope of the frequency spectrum.
Advantages:
Ideal for signals with a dominant frequency or periodicity.
Captures global periodic trends in the signal.
Efficient due to the FFT’s computational speed.
Graph Explanation:
Red dots correspond to detected peaks, largely highlighting strong frequency components.
Peaks appear consistent in rhythm but may miss finer oscillations or transients not aligned with the dominant frequency.
Transform-Based

## Slide 24
Short-Term Energy Peaks
Mechanism
Computes energy in short sliding windows using the squared amplitude sum.
Peaks are detected when energy exceeds a preset threshold.
Advantages:
Effective for transient detection and burst-like events.
Simple and fast to compute.
Useful in noisy or speech-like signals.
Graph Explanation:
Only one red peak is detected due to the energy threshold being too strict or low signal power across windows.
Misses periodic peaks due to lack of energy contrast.
Statistical/Energy

## Slide 25
RMS Energy Peaks
Mechanism
Calculates Root Mean Square (RMS) of signal segments over a sliding window.
Peaks are where RMS energy exceeds a threshold.
Advantages:
Captures overall signal power, suppressing random fluctuations.
Robust to noise and baseline shifts.
Graph Explanation:
Only a single peak is detected, similar to short-term energy.
Likely due to low RMS contrast in the signal; majority of frames don’t cross the set threshold.
Statistical/Energy

## Slide 26
RMS Energy Peaks
Mechanism
Calculates Root Mean Square (RMS) of signal segments over a sliding window.
Peaks are where RMS energy exceeds a threshold.
Advantages:
Captures overall signal power, suppressing random fluctuations.
Robust to noise and baseline shifts.
Graph Explanation:
Only a single peak is detected, similar to short-term energy.
Likely due to low RMS contrast in the signal; majority of frames don’t cross the set threshold.
Statistical/Energy

## Slide 27
Slope Sum Function (SSF)
Mechanism
Computes cumulative positive slopes over a moving window.
Peaks occur where this slope sum is maximal.
Advantages:
Effective in highlighting rising edges, useful in pulse-like signals.
Good temporal localization of rapid changes.
Graph Explanation:
Only a single early peak is detected, likely due to strong initial upward slope.
Later oscillations lack steep rises, so SSF remains low.
Gradient/SSF-Based

## Slide 28
First-Order Difference Histogram
Mechanism
Computes the histogram of first-order signal differences.
Peaks are detected where the histogram shows high frequency of rising edges.
Advantages:
Captures abrupt changes efficiently.
Suitable for signals with impulsive or sudden transitions.
Graph Explanation:
Detected peaks occur near sharp downward transitions in the early signal.
Later smoother sections lack strong first-order difference, hence no peaks found.
Gradient/SSF-Based

## Slide 29
Integrated Slope Method
Mechanism
Accumulates only positive slope changes over a window.
Peaks are where slope accumulation exceeds a threshold.
Advantages:
Enhances sharp upslopes, suppresses noise.
Effective for detecting steep upward peaks.
Graph Explanation:
Peaks align with sharp rising segments across the signal.
Consistent detection in both early and late oscillations.
Gradient/SSF-Based

## Slide 30
Automatic Multiscale Peak Detection
Mechanism
Scans signal using multiple window sizes
Peaks are local maxima across all scales, forming a peak consensus.
Advantages:
Fully automatic (no threshold/tuning).
Robust to noise and baseline drift.
Graph Explanation:
Detects only the most consistent peaks, ignoring smaller oscillations.
Selected peaks show high scale-invariance.
AMPD

## Slide 31
Modified AMPD
Mechanism
Builds on original AMPD by refining peak detection range.
Adds preprocessing (e.g., smoothing or bandpass filtering) before scale consensus.
Advantages:
Better noise suppression, fewer false peaks.
Captures all meaningful peaks, especially in rPPG signals.
Graph Explanation:
More responsive to local peaks compared to AMPD.
Enhanced detection of periodic peaks with good balance.
AMPD

## Slide 32
Fast AMPD
Mechanism
Optimization of standard AMPD using matrix sparsity and vectorization.
Reduces redundant computations in the local maxima matrix.
Advantages:
Much faster on long signals or large datasets.
Same accuracy as AMPD, but improved runtime.
Graph Explanation:
Peaks match Modified AMPD, but detected more efficiently.
No over-smoothing or peak suppression visible.
AMPD

## Slide 33
FAMD (Frequency-Adaptive Multiscale)
Mechanism
Combines multiscale peak detection with adaptive filtering based on signal’s dominant frequency.
Enhances local maxima that match expected periodicity.
Advantages:
Works well for quasi-periodic signals like physiological waveforms.
Reduces noise-induced false peaks.
Graph Explanation:
Only strong, periodic peaks are retained.
Sparse detection, focused on true signal cycles.
AMPD

## Slide 34
Statistical Threshold
Mechanism
Peaks are detected if the amplitude exceeds a fixed threshold.
Threshold is manually set or based on signal stats,
Advantages:
Simple and fast.
Effective when signal amplitude is consistent.
Graph Explanation:
Only higher peaks are detected.
May cause multiple detections near single peaks if not filtered.
Threshold-Based

## Slide 35
Adaptive Threshold
Mechanism
Threshold is dynamically adjusted over time using a sliding window (local mean or median).
Peaks are those exceeding local threshold.
Advantages:
Robust to baseline drift and amplitude changes.
Fewer false positives than static methods.
Graph Explanation:
Peaks detected based on local signal context, even if global amplitude varies.
Threshold-Based

## Slide 36
Max Threshold
Mechanism
A global maximum-based threshold is computed (60–90% of peak value).
All points above this value are marked as peaks.
Advantages:
Simple, fast, and no parameter tuning.
Works well when peak amplitudes are consistently high.
Graph Explanation:
Peaks are detected if they exceed a fixed high-level threshold, relative to the signal max.
Threshold-Based

## Slide 37
Bandpass + Local Maxima
Mechanism
Signal is filtered in a target frequency range
Local maxima are extracted from the filtered signal.
Advantages:
Removes noise outside target band.
Simple and effective for periodic signals.
Graph Explanation:
Peaks occur at maxima in the bandpassed waveform, aligning with physiological rhythms.
Filtering-Based

## Slide 38
Savitzky–Golay Peak Detection
Mechanism
Smooths signal using polynomial fitting over a sliding window.
Peaks are detected from the smoothed curve’s maxima.
Advantages:
Preserves peak shape and height.
Reduces noise without excessive distortion.
Graph Explanation:
Peaks lie at local maxima of the smoothed curve, improving detection in noisy segments.
Filtering-Based

## Slide 39
Morphological Gradient
Mechanism
Computes the difference between dilation and erosion of the signal (max - min filter).
Peaks are located at local intensity gradients.
Advantages:
Robust against noise.
Enhances edge like features in the signal.
Graph Explanation:
Peaks are aligned with sharp rises in the signal, revealed by strong morphological contrast.
Filtering-Based

## Slide 40
Gaussian Smoothing Peaks
Mechanism
Applies Gaussian filter to reduce noise.
Detects local maxima in the smoothed signal.
Advantages:
Suppresses high-frequency noise.
Improves peak shape clarity.
Graph Explanation:
Peaks appear smoother and more distinct after noise reduction by the Gaussian blur.
Filtering-Based

## Slide 41
Median Filtered Peaks
Mechanism
Applies median filter to remove spikes/noise.
Finds peaks in the smoothed signal.
Advantages:
Effective against impulsive noise.
Preserves edge features better than Gaussian.
Graph Explanation:
Median smoothing enhances peak visibility by removing sharp outliers before peak detection.
Filtering-Based

## Slide 42
HR Estimation: Peak Vally Detector
100 frames -> 3 sec
HR (Normal): 93.33333333333333
Difference with GT=> 3.6666666666666714
HR (Z-score filtered): 93.33333333333333
Difference with GT=> 3.6666666666666714
HR (Threshold filtered): 93.33333333333333
Difference with GT=> 3.6666666666666714
Average Ground Truth HR: 97.00 bpm (100 & 300 Frames)
300 frames -> 10 sec
HR (Normal): 101.67647058823532
Difference with GT=> 4.676470588235318
HR (Z-score filtered): 101.67647058823532
Difference with GT=> 4.676470588235318
HR (Threshold filtered): 101.67647058823532
Difference with GT=> 4.676470588235318
Threshold-Based Filtering
HR between 40 and 180 bpm means RR between 0.33 s to 1.5 s
Z-Score Filtering
|z| > threshold (commonly 2.0 or 3.0).

## Slide 43
HR Estimation: Adaptive Threshold
100 frames -> 3 sec
HR (Normal): 99.99999999999999
Difference with GT=> 2.999999999999986
HR (Z-score filtered): nan
Difference with GT=> nan
HR (Threshold filtered): 99.99999999999999
Difference with GT=> 2.999999999999986
100 frames -> 3 sec
HR (Normal): 75.8848674604977
Difference with GT=> 21.115132539502298
HR (Z-score filtered): 84.47829131652661
Difference with GT=> 12.521708683473392
HR (Threshold filtered): 84.47829131652661
Difference with GT=> 12.521708683473392
Average Ground Truth HR: 97.00 bpm (100 & 300 Frames)
Threshold-Based Filtering
HR between 40 and 180 bpm means RR between 0.33 s to 1.5 s
Z-Score Filtering
|z| > threshold (commonly 2.0 or 3.0).

## Slide 44
HR Estimation: Fast AMPD
100 Frames – 3sec
HR (Normal): 93.33333333333333
Difference with GT=> 3.6666666666666714
HR (Z-score filtered): 93.33333333333333
Difference with GT=> 3.6666666666666714
HR (Threshold filtered): 93.33333333333333
Difference with GT=> 3.6666666666666714
300 frames -> 3 sec
HR (Normal): 101.67647058823532
Difference with GT=> 4.676470588235318
HR (Z-score filtered): 101.67647058823532
Difference with GT=> 4.676470588235318
HR (Threshold filtered): 101.67647058823532
Difference with GT=> 4.676470588235318
Average Ground Truth HR: 97.00 bpm (100 & 300 Frames)
Threshold-Based Filtering
HR between 40 and 180 bpm means RR between 0.33 s to 1.5 s
Z-Score Filtering
|z| > threshold (commonly 2.0 or 3.0).

## Slide 45
HR Estimation: savgol_peak
100 Frames – 3sec
HR (Normal): 139.0909090909091
Difference with GT=> 42.09090909090909
HR (Z-score filtered): 139.0909090909091
Difference with GT=> 42.09090909090909
HR (Threshold filtered): 117.61363636363637
Difference with GT=> 20.613636363636374
300 frames -> 3 sec
HR (Normal): 127.9846774060032
Difference with GT=> 30.9846774060032
HR (Z-score filtered): 117.23371974387841
Difference with GT=> 20.23371974387841
HR (Threshold filtered): 110.04930106013698
Difference with GT=> 13.049301060136983
Average Ground Truth HR: 97.00 bpm (100 & 300 Frames)
Threshold-Based Filtering
HR between 40 and 180 bpm means RR between 0.33 s to 1.5 s
Z-Score Filtering
|z| > threshold (commonly 2.0 or 3.0).

## Slide 46
Peak Finding Algorithms
Peak Finding Agorithm
Local Maxima
Envelope-Based
Transform-Based
Statistical/Energy
Gradient/SSF-Based
AMPD
Threshold-Based
Filtering-Based
Pan-Tompkins
Algorithm
Squared Derivative
Method
Differential
Zero-Crossing
Standard Local
Maxima
Sliding Window
Maxima
Peak-Valley Detector
Hilbert Envelope Peaks
Moving Average
Envelope
Autocorrelation-Based Peaks
Continuous Wavelet Transform (CWT) Peaks
FFT Envelope Peaks
Short-Term Energy
Peaks
RMS Energy Peaks
Spectral Entropy
Gradient
Slope Sum Function
(SSF)
First-Order Difference Histogram
Integrated Slope Method
Derivative-Based
Zero-Crossing
Second Derivative
Zero-Crossing
Automatic Multiscale Peak Detection
Modified AMPD
Fast AMPD
FAMD (Frequency-Adaptive Multiscale)
Statistical Threshold
Adaptive Threshold
Max Threshold
Morphological
Gradient
Gaussian Smoothing
Peaks
Median Filtered Peaks
Bandpass +
Local Maxima
Savitzky–Golay Peak
Detection