## Slide 1
Rafsan Jany
Korean Convergence Medical Science
Paper Review and Phase Alignment

## Slide 2
1
Study - 4
Paper Name: The role of face regions in remote photoplethysmography for contactless heart rate monitoring
Authors: Maksym Bondarenko, Carlo Menon & Mohamed Elgendi (2025)
Description:
To systematically review 70 studies and analyze how the choice and number of Regions of Interest (ROIs: hole face, forehead, cheeks, nose, chin, multiple patches) influence the accuracy of rPPG-based heart rate (HR) estimation.
Methods:
Excluded: non-RGB, NIR/IR, theoretical, non-facial or contact-based methods.
Datasets analyzed: 20 public sets (UBFC-rPPG, PURE, COHFACE, VIPL-HR, MAHNOB-HCI).
Metrics: MAE (Mean Absolute Error) and RMSE (Root Mean Square Error).
ROI Selection
Most common: holistic (whole) face (36.8% of studies).
Best performing: forehead (24.5%) and cheeks (21.7%), due to large flat skin surfaces, stable perfusion, low hair/accessory interference .
Nose and chin were less reliable (6.6% and 4.7%).
ROI Quantity
More ROIs generally improved performance, but diminishing returns beyond ~60 patches.
Best results clustered around 3 ROIs and ~60 ROIs (LOESS analysis, Fig. 4).
Too many patches (>100) increased computation and sometimes added noise.

## Slide 3
2
Study - 4
Paper Name: The role of face regions in remote photoplethysmography for contactless heart rate monitoring
Authors: Maksym Bondarenko, Carlo Menon & Mohamed Elgendi (2025)
ROI Selection
Most common: holistic (whole) face (36.8% of studies).
Best performing: forehead (24.5%) and cheeks (21.7%), due to large flat skin surfaces, stable perfusion, low hair/accessory interference .
Nose and chin were less reliable (6.6% and 4.7%).

## Slide 4
3
Study - 4
Paper Name: The role of face regions in remote photoplethysmography for contactless heart rate monitoring
Authors: Maksym Bondarenko, Carlo Menon & Mohamed Elgendi (2025)
ROI Quantity
More ROIs generally improved performance, but diminishing returns beyond ~60 patches.
Best results clustered around 3 ROIs and ~60 ROIs.
Too many patches (>100) increased computation and sometimes added noise.

## Slide 5
4
Study - 4
Paper Name: The role of face regions in remote photoplethysmography for contactless heart rate monitoring
Authors: Maksym Bondarenko, Carlo Menon & Mohamed Elgendi (2025)
Performance Results
Algorithm / Model | UBFC MAE / RMSE | PURE MAE / RMSE | COHFACE MAE / RMSE | VIPL-HR MAE / RMSE | MAHNOB-HCI MAE / RMSE
MAR-rPPG (Qian et al. 2024) | 0.17 / 0.41 | 0.37 / 0.68 | – | 4.36 / 6.92 | –
TokenLearner rPPG (Zhao et al. 2024) | 0.12 / 0.35 | 0.08 / 0.29 | – | – | –
Dual-GAN rPPG (Hao Lu et al. 2021) | 0.44 / 0.67 | 0.82 / 1.31 | – | 4.93 / 7.68 | –
Graph-based rPPGNet (Si-Qi Liu et al. 2024) | 0.31 / 0.98 | 0.18 / 0.41 | 0.64 / 1.89 | 3.13 / 3.97 | –
Traditional ROI + ML Fusion (Yaran Duan et al. 2023) | 2.78 / 3.35 | 2.76 / 4.15 | 2.77 / 3.87 | 3.45 / 4.17 | –

## Slide 6
5
Study - 4
Paper Name: The role of face regions in remote photoplethysmography for contactless heart rate monitoring
Authors: Maksym Bondarenko, Carlo Menon & Mohamed Elgendi (2025)
“3 patches” vs. “60 patches”
3 patches = only a few large ROIs (forehead + 2 cheeks) Advantage: simple, fast, still accurate. Good enough for many datasets (PURE, UBFC).
60 patches = face divided into many small sub-ROIs (triangulation into dozens of segments). Advantage: can adaptively choose the best-quality patches (exclude noisy/occluded ones). Disadvantage: computationally heavy, diminishing returns after ~60 patches.
Which areas do those 60 patches usually cover?
Researchers don’t pick 60 unique best patches. Instead:
1.They subdivide the whole face into many small triangular or rectangular regions. Example: Delaunay triangulation (Gao et al.) created up to 898 micro-ROIs. Then only the best ~30–60 are selected (based on signal quality index, SNR, etc.).
2.These patches are concentrated in the forehead and cheek regions because: Flat, hair-free, less occlusion, good blood perfusion. Symmetry helps with fallback if one side is blocked.

## Slide 7
6
Study – 1
Paper Name: Continual Learning for Remote Physiological Measurement: Minimize Forgetting and Simplify Inference
Authors: Qian Liang, Yan Chen, and Yang Hu
Description:
This study introduces ADDP, a rehearsal-free continual learning framework for remote photoplethysmography (rPPG). It addresses catastrophic forgetting when models are trained sequentially on different datasets, by combining adapter-based fine-tuning, domain prototypes, and prototype-based augmentation/inference simplification to improve cross-domain robustness
Datasets Used:
The authors evaluate on five benchmark datasets under a new domain incremental learning (DIL) protocol:
VIPL-HR – large, multi-scenario dataset (used as initial training base).
PURE – six activities (sitting, talking, head movements).
UBFC-rPPG – uncompressed indoor/outdoor lighting variations.
BUAA-MIHR – low- to medium-light recordings.
MMPD – diverse skin tones (Fitzpatrick 3–6), multiple lighting and activity settings.
Methods Used
Adapter-based Fine-tuning -> Freeze backbone (Uniformer + DiffNorm), only update lightweight adapters.
Domain Prototypes -> Extract style prototypes (illumination, skin tone) and noise prototypes (motion, compression artifacts).
Prototype-based Augmentation -> Replay “pseudo-old” samples by mixing current training data with past domain prototypes.
Prototype-based Inference Simplification -> During testing, map new samples into the style the model knows best (lowest MAE prototype).

## Slide 8
7
Study – 1
Paper Name: Continual Learning for Remote Physiological Measurement: Minimize Forgetting and Simplify Inference
Authors: Qian Liang, Yan Chen, and Yang Hu
Model Description
The authors Backbone: Uniformer + DiffNorm module Uniformer captures both local (3D conv) and global (self-attention) features. DiffNorm fuses appearance + dynamic differences, boosting HR signal quality.
Why better: Unlike ViT (fails on rPPG) and PhysFormer (good but limited), Uniformer + DiffNorm captures subtle color changes. Adapter finetuning avoids catastrophic forgetting while remaining lightweight. Domain prototypes allow style/condition adaptation without storing raw data (privacy-friendly).
Unique Contributions
First exploration of continual learning (domain incremental learning) in rPPG.
Introduces ADDP: rehearsal-free, privacy-preserving, efficient adaptation across datasets.
Proposes DiffNorm + Uniformer backbone specifically tailored for rPPG signals.
Establishes a benchmark protocol for rPPG DIL, enabling fair future comparisons.

## Slide 9
8
Study – 1
Paper Name: Continual Learning for Remote Physiological Measurement: Minimize Forgetting and Simplify Inference
Authors: Qian Liang, Yan Chen, and Yang Hu
Continual Learning (CL)
Normally, deep learning models are trained once on a fixed dataset → then tested.
In the real world, data arrives sequentially (new environments, subjects, cameras, lighting, activities).
Continual Learning (CL) = training a model step-by-step on new tasks/domains without forgetting what it learned earlier.
The challenge: avoid catastrophic forgetting (where learning new data overwrites past knowledge).
Incremental Learning (DIL)
A type of continual learning where: The task stays the same (predicting HR from face videos). But the domain keeps changing (different datasets: PURE → UBFC → MMPD → VIPL-HR).
Each dataset has different lighting, motion, skin tones, camera quality = new “domain.”
Goal: as the model sees each domain sequentially, it should: Learn to generalize to the new one. Retain performance on the old ones

## Slide 10
9
Study – 1
Paper Name: Continual Learning for Remote Physiological Measurement: Minimize Forgetting and Simplify Inference
Authors: Qian Liang, Yan Chen, and Yang Hu
Result Table
Methods | Training Type / Meaning | StdN ↓ | MAEN ↓ | RMSEN ↓ | RN ↑
Upper Bound (joint training) | Joint training on all datasets at once (ideal best case, no forgetting) | 5.25 ± 0.09 | 3.45 ± 0.09 | 5.31 ± 0.24 | 0.85 ± 0.01
Lower Bound (naïve finetune) | Sequential fine-tuning without protection → severe catastrophic forgetting | 8.73 ± 0.79 | 6.54 ± 0.27 | 9.46 ± 0.49 | 0.69 ± 0.02
EWC [25] | Regularization – penalizes changes to important weights (Elastic Weight Consolidation) | 6.88 ± 0.89 | 4.51 ± 0.33 | 7.02 ± 0.86 | 0.76 ± 0.04
ANCL-EWC [23] | Auxiliary network + EWC – adds small side net to balance stability vs. plasticity | 6.87 ± 0.36 | 4.39 ± 0.08 | 6.97 ± 0.36 | 0.77 ± 0.02
LwF [29] | Distillation – “Learning without Forgetting,” matches outputs of old model | 6.59 ± 0.14 | 4.38 ± 0.10 | 6.30 ± 0.67 | 0.78 ± 0.02
ANCL-LwF [23] | Auxiliary network + Distillation – improves LwF with extra side net | 6.37 ± 0.28 | 4.13 ± 0.13 | 6.12 ± 0.25 | 0.80 ± 0.01
LAE-online [14] | Low-rank Adapter, online – parameter-efficient fine-tuning (adapters updated per task) | 6.97 ± 0.56 | 4.70 ± 0.38 | 7.13 ± 0.56 | 0.78 ± 0.03
LAE-offline [14] | Low-rank Adapter, offline – stores adapters for later use, but fails in regression | 12.06 ± 0.86 | 7.94 ± 0.35 | 12.36 ± 0.96 | 0.59 ± 0.04
LAE-ensemble [14] | Adapter ensemble – combines multiple adapters → heavy, poor for rPPG regression | 12.37 ± 0.42 | 8.07 ± 0.24 | 12.61 ± 0.51 | 0.55 ± 0.01
S-Prompts [52] | Soft prompts / adapter pool – lightweight prompts for each task, no rehearsal needed | 5.77 ± 0.18 | 3.85 ± 0.05 | 5.82 ± 0.19 | 0.81 ± 0.01
ADDP (ours) | Adapters + Domain Prototypes – rehearsal-free, style/noise prototypes + inference map | 5.56 ± 0.04 | 3.70 ± 0.07 | 5.59 ± 0.02 | 0.83 ± 0.01

## Slide 11
10
Study – 1
Paper Name: Continual Learning for Remote Physiological Measurement: Minimize Forgetting and Simplify Inference
Authors: Qian Liang, Yan Chen, and Yang Hu
Result Table
Interpretation:
Style prototypes alone reduce error → model remembers illumination/skin tone variations.
Adding noise prototypes further improves robustness to motion/compression.
Inference simplification (Sim.) gives the largest gain, letting the model handle new tasks by mapping them to familiar styles.
Their full ADDP (Style + Noise + Sim.) achieves the best performance.
Style | Noise | Sim. | StdN ↓ | MAEN ↓ | RMSEN ↓ | RN ↑
✗ | ✗ | ✗ | 7.06 ± 0.90 | 4.59 ± 0.53 | 7.24 ± 0.96 | 0.78 ± 0.03
✓ | ✗ | ✗ | 6.68 ± 0.67 | 4.23 ± 0.21 | 6.83 ± 0.60 | 0.73 ± 0.06
✓ | ✓ | ✗ | 6.66 ± 0.61 | 4.10 ± 0.20 | 6.61 ± 0.63 | 0.75 ± 0.05
✓ | ✓ | TS | 5.71 ± 0.01 | 3.91 ± 0.04 | 5.89 ± 0.03 | 0.81 ± 0.00
✓ | ✓ | ✓ | 5.56 ± 0.04 | 3.70 ± 0.07 | 5.59 ± 0.02 | 0.83 ± 0.01

## Slide 12
11
Study – 2
Paper Name: Dual-path TokenLearner for Remote Photoplethysmography-based Physiological
Measurement with Facial Videos
Authors: Wei Qian, Dan Guo∗, Kun Li, Xilan Tian, Meng Wang
Description:
The paper proposes Dual-TL, a Transformer-based rPPG model that learns two global tokens in parallel—Spatial TokenLearner (S-TL) across facial ROI combinations and Temporal TokenLearner (T-TL) across frames—to suppress noise (motion, illumination, occlusion) and predict rPPG signals/HR robustly. It achieves state-of-the-art results on multiple benchmarks and strong cross-dataset generalization.
Datasets Used:
VIPL-HR – large, multi-scenario dataset (used as initial training base).
PURE – six activities (sitting, talking, head movements).
UBFC-rPPG – uncompressed indoor/outdoor lighting variations.
MMSE-HR (used for cross-dataset testing)
Methods Used
Adapter-based Fine-tuning -> Freeze backbone (Uniformer + DiffNorm), only update lightweight adapters.
Domain Prototypes -> Extract style prototypes (illumination, skin tone) and noise prototypes (motion, compression artifacts).
Prototype-based Augmentation -> Replay “pseudo-old” samples by mixing current training data with past domain prototypes.
Prototype-based Inference Simplification -> During testing, map new samples into the style the model knows best (lowest MAE prototype).

## Slide 13
12
Study – 2
Paper Name: Dual-path TokenLearner for Remote Photoplethysmography-based Physiological
Measurement with Facial Videos
Authors: Wei Qian, Dan Guo∗, Kun Li, Xilan Tian, Meng Wang
Methods Used
Feature prep (MSTmap): Detect landmarks with OpenFace; build 63 ROI combinations; average RGB/YUV; stack into N×C×T spatiotemporal map with max–min normalization. -> Freeze backbone (Uniformer + DiffNorm), only update lightweight adapters.
Dual TokenLearners: S-TL models dependencies across ROI combinations (spatial sequence). T-TL models heartbeat periodicity across frames (temporal sequence).
Fusion & head: Concatenate spatial/temporal tokens → FFNreg (GELU) → rPPG sequence (length T). Train with Negative Pearson loss; convert to HR via bandpass (0.75–2.5 Hz) + peak/FFT.
Design choices validated: YUV color space best; tokens crucial (large drops without them).

## Slide 14
13
Study – 2
Paper Name: Dual-path TokenLearner for Remote Photoplethysmography-based Physiological
Measurement with Facial Videos
Authors: Wei Qian, Dan Guo∗, Kun Li, Xilan Tian, Meng Wang
The authors don’t feed raw face video directly into the model. Instead, they preprocess the video into a structured input called an MSTmap (Multi-Spectral Temporal map).
Instead of just forehead/cheeks, they combine different face parts into 63 ROI groups (like small patches covering different skin areas).
For each ROI, they compute the average RGB/YUV values over time.
Y = Luminance (brightness)
How light or dark the pixel is.
Similar to a grayscale version of the image.
U = Chrominance (blue projection)
Encodes how much blue is in the color relative to brightness.
V = Chrominance (red projection)
Encodes how much red is in the color relative to brightness.
# Conversion matrix (BT.601 standard)
M = np.array([
[0.299, 0.587, 0.114 ],
[-0.14713, -0.28886, 0.436 ],
[0.615, -0.51499, -0.10001 ]
])

## Slide 15
14
Study – 2
Model Input: MSTmap (63 × 3 × T). Model Size: 4M param approximate.
Time (T frames {10 sec window}) ───────────────►
┌──────────────────────┐
ROI #1 │ [Y1, U1, V1] [Y2, U2, V2] ... [YT, UT, VT] │
ROI #2 │ [Y1, U1, V1] [Y2, U2, V2] ... [YT, UT, VT] │
ROI #3 │ [Y1, U1, V1] [Y2, U2, V2] ... [YT, UT, VT] │
... │ ... │
ROI #63 │ [Y1, U1, V1] [Y2, U2, V2] ... [YT, UT, VT] │
└──────────────────────┘
↑
3 color channels (YUV)

## Slide 16
15
Study – 2
Train → Test | Method | MAE ↓ | RMSE ↓ | r ↑
VIPL-HR → MMSE-HR | CHROM [2] | 3.81 | 6.42 | 0.89
 | POS [3] | 3.55 | 6.19 | 0.90
 | CVD [14] | 3.12 | 6.08 | 0.91
 | PhysFormer [21] | 2.84 | 5.36 | 0.92
 | Dual-TL (ours) | 2.25 | 5.11 | 0.93
PURE → MMSE-HR | CHROM [2] | 4.10 | 6.94 | 0.85
 | POS [3] | 3.99 | 6.84 | 0.86
 | CVD [14] | 3.72 | 6.80 | 0.87
 | rPPGNet [16] | 3.43 | 6.62 | 0.88
 | PhysFormer [21] | 2.84 | 5.89 | 0.90
 | Dual-TL (ours) | 2.51 | 5.27 | 0.93
MMSE-HR → PURE | CHROM [2] | 3.68 | 6.04 | 0.87
 | POS [3] | 3.53 | 5.92 | 0.88
 | CVD [14] | 2.26 | 3.57 | 0.96
 | PhysFormer [21] | 2.41 | 3.61 | 0.95
 | Dual-TL (ours) | 2.18 | 3.23 | 0.98
UBFC-rPPG → PURE | PhysNet [11] | 31.45 | 39.25 | –0.01
 | PhysFormer [21] | 23.63 | 30.70 | 0.07
 | Dual-TL (ours) | 14.12 | 23.47 | 0.23

## Slide 17
16
Study – 3
Paper Name: Dual-path TokenLearner for Remote Photoplethysmography-based Physiological
Measurement with Facial Videos
Authors: Wei Qian, Dan Guo∗, Kun Li, Xilan Tian, Meng Wang
Description:
The paper proposes Dual-GAN, a framework with two adversarial networks that jointly model the blood volume pulse (BVP) signal and noise distribution. It also introduces a ROI Alignment & Fusion (ROI-AF) block to handle ROI-specific noise variations. This approach improves robustness for HR, HRV, and RF estimation in noisy, real-world conditions.
Datasets Used:
VIPL-HR – large, multi-scenario dataset (used as initial training base).
PURE – six activities (sitting, talking, head movements).
UBFC-rPPG – uncompressed indoor/outdoor lighting variations.
Methods Used
Spatial-Temporal Map (STMap): Constructed from face ROI color time series as input.
Dual-GAN architecture: BVP-GAN: Learns to predict clean BVP signals from STMap. Noise-GAN: Learns to model realistic noise distributions and generate synthetic STMaps with noise. Shared discriminator links the two GANs.
ROI-AF Block: Performs ROI-wise convolution + channel attention. Reduces inconsistency between ROIs and enlarges receptive field.
Losses: Negative Pearson correlation, frequency-domain loss, HR regression (auxiliary task).

## Slide 18
17
Study – 3
Paper Name: Dual-path TokenLearner for Remote Photoplethysmography-based Physiological
Measurement with Facial Videos
Authors: Wei Qian, Dan Guo∗, Kun Li, Xilan Tian, Meng Wang
Dataset / Setting | Method | MAE ↓ | RMSE ↓
UBFC-rPPG (intra) | Green | 2.05 | 2.93
 | CHROM | 1.67 | 2.58
 | POS | 1.62 | 2.51
 | CVD | 1.21 | 1.80
 | RhythmNet | 1.18 | 1.68
 | PulseGAN | 1.19 | 1.74
 | Dual-GAN (ours) | 0.44 | 0.67
PURE (intra) | Green | 1.94 | 3.18
 | CHROM | 1.72 | 2.95
 | POS | 1.61 | 2.82
 | HR-CNN | 1.03 | 1.74
 | 2SR | 0.92 | 1.59
 | Dual-GAN (ours) | 0.82 | 1.31
VIPL-HR (intra) | Green | 18.77 | 24.24
 | CHROM | 17.72 | 22.81
 | POS | 16.49 | 21.20
 | CVD | 12.80 | 16.20
 | DeepPhys | 8.94 | 12.24
 | RhythmNet | 7.33 | 10.97
 | Dual-GAN (ours) | 4.93 | 7.68
Cross-dataset (PURE→UBFC) | PulseGAN | 2.09 | 2.63
 | Dual-GAN (ours) | 0.74 | 1.02

## Slide 19
18
Study – 4
Paper Name: Efficient Remote Photoplethysmography with Temporal Derivative Modules and Time-Shift Invariant Loss
Authors: Joaquim Comas1, Adri`a Ruiz2, Federico Sukno
Description:
The paper proposes Dual-GAN, a framework with two adversarial networks that jointly model the blood volume pulse (BVP) signal and noise distribution. It also introduces a ROI Alignment & Fusion (ROI-AF) block to handle ROI-specific noise variations. This approach improves robustness for HR, HRV, and RF estimation in noisy, real-world conditions.
Datasets Used:
PURE
UBFC-rPPG
Methods Used
Input & preprocessing: face via MTCNN, enlarge box (+50%), resize to 128×128, downsample ground-truth PPG to 30 Hz; train on 256-frame clips (~8.5 s).
Spatial encoder (tiny): two 2D conv (3×3) + tanh + BN + avg-pool; outputs per-frame spatial features.
Temporal Derivative Module (TDM): chain of fixed 1D derivative filters over time (kernel [–2, –1, 0, 1, 2]), stacked as first- and second-order (DTC1, DTC2), then concat → 1×1 conv to produce the rPPG sequence. Emulates Taylor-series dynamics with skip connections.
TALOS loss (time-shift invariant): learns a per-subject discrete distribution pθ(k|s) over temporal offsets k; minimizes a mixture of MSEs between prediction and shifted ground truth, making training robust to PTT / desync.

## Slide 20
19
Study – 4
Paper Name: Efficient Remote Photoplethysmography with Temporal Derivative Modules and Time-Shift Invariant Loss
Authors: Joaquim Comas1, Adri`a Ruiz2, Federico Sukno
Method | PURE (RMSE↓ / MAE↓ / r↑) | UBFC (RMSE↓ / MAE↓ / r↑)
CHROM | 2.50 / 2.07 / 0.99 | 4.61 / 3.44 / 0.97
POS | 10.57 / 3.14 / 0.95 | 6.61 / 2.44 / 0.94
HR-CNN | 2.37 / 1.84 / 0.98 | –
DeepPhys | 2.31 / 1.84 / 0.99 | 3.63 / 2.90 / 0.98
PhysNet | 2.70 / 2.16 / 0.99 | 3.67 / 2.95 / 0.98
Zhan et al. | 2.29 / 1.82 / 0.99 | 3.17 / 2.44 / 0.99
Ours (TDM) | 2.33 / 1.89 / 0.99 | 3.31 / 2.54 / 0.99
Ours + TALOS | 2.30 / 1.83 / 0.99 | 3.08 / 2.32 / 0.99

## Slide 21
20
Study – 4
Paper Name: Efficient Remote Photoplethysmography with Temporal Derivative Modules and Time-Shift Invariant Loss
Authors: Joaquim Comas1, Adri`a Ruiz2, Federico Sukno
The network learns a proxy signal:
Strong enough periodicity to estimate HR accurately.
But not guaranteed to match real BVP morphology.

## Slide 22
21
Study – 5
Paper Name: Region of Interest Analysis Using Delaunay Triangulation for Facial Video-Based
Heart Rate Estimation
Authors: Haoyuan Gao, Chao Zhang, Shengbing Pei and Xiaopei Wu
Description:
This paper investigates how different facial regions contribute to rPPG-based heart rate (HR) estimation. By applying Delaunay triangulation to MediaPipe’s 468 landmarks, the face is divided into 898 triangular patches (ROIs), and their effectiveness for HR extraction is systematically analyzed. A data-driven ROI (DD-ROI) method is proposed, which significantly improves performance across datasets.
Datasets Used:
IIP-F (fixed webcam, 30 fps, 16M/10F, rest, talk, head rotation)
IIP-W (head-mounted webcam, 25 fps, 16M/10F, same tasks)
COHFACE (40 subjects, 20 fps, webcam, PPG reference)
PURE (10 subjects, 6 conditions, 30 fps, PPG reference)
UBFC-rPPG (50 videos, 30 fps, CMS50E PPG reference)
MAHNOB-HCI (27 subjects, 61 fps, multimodal incl. ECG)
Methods Used
Landmark detection – 468 face points via MediaPipe.
Delaunay triangulation – face partitioned into 898 triangular ROIs.
Signal extraction – RGB averaging per ROI → green-channel filtering (0.6–4 Hz).
HR estimation – via FFT peak frequency per 10s sliding window.
DD-ROI selection – sort ROIs by HR MAE; select top nr ROIs (≈40).
Comparison with classic ROIs (full face, 60% center, skin segmentation, cheeks).
Validation with multiple rPPG methods: CHROM, POS, GREEN, ICA, LGI, OMIT, PBV.

## Slide 23
22
Study – 5
Paper Name: Region of Interest Analysis Using Delaunay Triangulation for Facial Video-Based
Heart Rate Estimation
Authors: Haoyuan Gao, Chao Zhang, Shengbing Pei and Xiaopei Wu
Actually How It Works:
Facial video (frames)
↓
Face landmarks (468 points, MediaPipe)
↓
Delaunay triangulation → ~898 triangular ROIs
↓
For each frame & each ROI:
- Collect pixels inside ROI
- Compute mean R, G, B
↓
Build time series → R(t), G(t), B(t) per ROI
↓
Apply bandpass filter (0.6–4 Hz)
↓
Final output: ~898 filtered RGB signals

## Slide 24
23
Study – 5
Paper Name: Region of Interest Analysis Using Delaunay Triangulation for Facial Video-Based
Heart Rate Estimation
Authors: Haoyuan Gao, Chao Zhang, Shengbing Pei and Xiaopei Wu
~898 filtered RGB signals (from Step 3)
↓
For each ROI patch:
- Select a channel (often G(t) or combined RGB)
- Apply FFT (Fast Fourier Transform)
- Find dominant frequency peak
- Convert frequency → HR in BPM
↓
Compare with ground-truth HR (from PPG/ECG)
↓
Compute error = |HR_pred - HR_GT|
↓
Average across windows → MAE per ROI
↓
Final output: MAE score for each of the 898 ROIs

## Slide 25
24
Study – 5
Condition | Dataset(s) | Best ROI Method (MAE bpm) | Comparison (bpm)/ Notes
Intra-dataset (Resting) | PURE | DD-ROI: 1.70 | vs Face Det: 5.22, Skin Det: 2.25, Landmark: 2.37
 | IIP-W | DD-ROI: 2.24 | vs Face Det: 7.60, Skin Det: 5.76, Landmark: 4.99
 | IIP-F | DD-ROI: 6.19 | vs Face Det: 14.88, Skin Det: 9.09, Landmark: 8.18
Intra-dataset (Talking) | PURE | DD-ROI: 7.21 | Robust vs noise; Face Det: 12.72
 | IIP-W | DD-ROI: 7.22 | Best under head-mounted cam motion
 | IIP-F | DD-ROI: 9.04 | Still lowest error despite occlusion/noise
Cross-dataset (Trained PURE → Tested) | UBFC | DD-ROI: 19.40 | vs Face Det: 25.51, Skin Det: 22.29, Landmark: 22.20
 | COHFACE | DD-ROI: 7.41 | vs Face Det: 10.23, Skin Det: 8.29, Landmark: 8.08
Across rPPG Methods (PURE) | POS | DD-ROI: 1.70 | vs 2.25–2.62 for other ROI methods
 | CHROM | DD-ROI: 1.97 | vs 2.62 (Skin/Face)
 | GREEN | DD-ROI: 2.18 | vs 6.14 (Face Det), 2.78 (Skin)
 | ICA | DD-ROI: 3.19 | vs 9.88 (Face Det), 4.37 (Skin)
 | LGI | DD-ROI: 2.53 | vs 8.12 (Face Det), 3.65 (Skin)
 | OMIT | DD-ROI: 2.40 | vs 7.13 (Face Det), 3.12 (Skin)
 | PBV | DD-ROI: 2.16 | vs 6.02 (Face Det), 2.76 (Skin)

## Slide 26
25
FPS Based Phase Lag Fix
fps_affine
tau = a + b / fps
a = a small constant delay (0.10 s)
b / fps = “about b frames of delay”, converted to seconds (b=12 → 12/fps)
SHIFT_INVARIANT_LOSS
For each batch, it rolls the GT by ±k samples (say k=8), computes MAE for each shift, and keeps the minimum.
Then convert to samples and roll the BVP left by that many samples:
s0 = round(tau * fps)
BVP <- roll_left(BVP, s0)
Together: why it works
fps_affine: removes the big, predictable part of the lag (dataset/system-level).
SI loss: makes the model robust to the small, unpredictable part of the lag (subject/window-level).
Raw GT BVP (delayed) Predicted BVP
|-----------Δ-------------->|
fps_affine
(coarse shift)
|
v
Approx. aligned GT -------->|
(lag mostly removed)
SI Loss
(±k sample tolerance)
|
v
Final loss computed without punishing tiny misalignments

## Slide 27
26
FPS Based Phase Lag Fix

## Slide 28
27
Experiment on Window Acceptance Based on Head Position
Lets Move to CODE for LIVE DEMONSTRATION

## Slide 29
28
Beat-to-Beat Variability from PPG: HRV and Pulse Amplitude
Heart Rate Variability (HRV)
PPG Amplitude Variability (PAV)
Amplitude
Stress monitoring (sympathetic activation) ↑ Sympathetic drive → ↓ HRV (↓ RMSSD, ↓ HF power) ∝ ↓ Pulse amplitude, ↑ Beat-to-beat amplitude lability
When the sympathetic system ramps up (stress, arousal), parasympathetic activity drops, leading to lower HRV (lower RMSSD and HF power), while at the same time pulse amplitude decreases and becomes more unstable from beat to beat.
Peripheral circulation assessment ↓ Mean pulse amplitude, ↑ Amplitude variability ∝ ↓/unstable peripheral perfusion & microvascular health (site, temp, vasoactive state)
If low average pulse amplitude plus unstable fluctuations beat-to-beat, it’s a signal of compromised or unstable peripheral circulation, reflecting issues in microvascular health and perfusion dynamics.
Anesthesia monitoring (vascular tone & hemodynamics) Amplitude level & resp.-linked variation ∝ vasodilation / vasoconstriction & fluid responsiveness → basis of pleth indices (PVI, SPI)
Both the absolute amplitude and its breathing-linked variability are physiological markers of vascular tone and fluid status.

## Slide 30
29
Beat-to-Beat Variability from PPG: HRV and Pulse Amplitude
Metric | Predicted | Ground Truth
Beats (N+1) | 16 | 16
Mean IBI (s) | 0.493 | 0.498
Mean HR (bpm) | 121.622 | 120.536
SDNN (s) | 0.023 | 0.015
RMSSD (s) | 0.028 | 0.020
Metric | Predicted | Ground Truth
Beats (M) | 15 | 15
Mean amplitude (a.u.) | 2.525 | 2.766
SD amplitude (a.u.) | 0.139 | 0.097
CV amplitude (%) | 5.506 | 3.510
RMS ΔA (a.u.) | 0.226 | 0.099
Pred. BVP with peaks (red), preceding troughs (green), and per-beat amplitudes (brown).
Ground-truth BVP with peaks, troughs, and amplitudes.
Pred. BVP with IBI (s) and instantaneous HR (bpm).
Ground-truth BVP with IBI and instantaneous HR.

## Slide 31
30
Beat-to-Beat Variability from PPG: HRV and Pulse Amplitude
UBFC 8 s window: Prediction (top) and Ground Truth (bottom) with the corresponding low-frequency baseline
(DC) extracted via zero-phase low-pass filtering at
f_c=0.4 Hz (Sec. 6.1). Both traces exhibit a coherent
respiratory oscillation in the baseline,
yielding the same respiratory-rate estimate of 15 brpm. Identical preprocessing was applied to Prediction and
Ground Truth.
Metric | Pred | GT | ¦Δ¦
RR (brpm) | 15.000 | 15.000 | 0.000
Resp. SNR (dB) | 2.345 | 4.943 | 2.598
RPF (0.10–0.40 Hz) | 0.000 | 0.000 | 0.000
Baseline peak-to-peak swing | 0.314 | 0.602 | 0.288
Metric | Pred | GT | ¦Δ¦
Beats (for amplitudes) | 16 | 16 | 0
Mean amplitude (a.u.) | 2.519 | 2.749 | 0.230
SD amplitude (a.u.) | 0.137 | 0.095 | 0.042
CV amplitude (%) | 5.443 | 3.467 | 1.976
Respiratory Features
Pulse Amplitude Variability

## Slide 32
31
Beat-to-Beat Variability from PPG: HRV and Pulse Amplitude
UBFC 8 s window: Prediction (top) and Ground Truth (bottom) with the corresponding low-frequency baseline
(DC) extracted via zero-phase low-pass filtering at
f_c=0.4 Hz (Sec. 6.1). Both traces exhibit a coherent
respiratory oscillation in the baseline,
yielding the same respiratory-rate estimate of 15 brpm. Identical preprocessing was applied to Prediction and
Ground Truth.
Metric | Pred | GT | ¦Δ¦
RR (brpm) | 15.000 | 15.000 | 0.000
Resp. SNR (dB) | 2.345 | 4.943 | 2.598
RPF (0.10–0.40 Hz) | 0.000 | 0.000 | 0.000
Baseline peak-to-peak swing | 0.314 | 0.602 | 0.288
Metric | Pred | GT | ¦Δ¦
Beats (for amplitudes) | 16 | 16 | 0
Mean amplitude (a.u.) | 2.519 | 2.749 | 0.230
SD amplitude (a.u.) | 0.137 | 0.095 | 0.042
CV amplitude (%) | 5.443 | 3.467 | 1.976
Respiratory Features
Pulse Amplitude Variability